RP Maintenance

Summary

The AccuTOF user should take care of RP maintenance.

This rotary pump (LC-MS interface) is pumping lower vacuum area then the GC-MS system.

Accordingly, the user should check oil level regularly.

Check the mist filter reservoir oil level and drain back into the RP using the Oil Return drain valve.

If necessary, add oil to the rotary pump.

Also, we recommend changing the rotary pump oil and mist filter once every 6 to 12 months.

# Check mist filter

(Check once per 1-2 weeks)

Check Mist Filter level

Oil Return drain valve

Open valve until oil goes back to

the RP.

(Oil level in the filter tank will drop,

approximately: 1-2 minutes)

# Open the valve only while in the

“Evacuation Ready” mode.

Add oil

- Open the front door of the AccuTOF to check oil level.

- If oil level is close to the lower level, oil should be added.

2-1  Open right side panel of AccuTOF.

2-2  Vent evacuation (Do vent process)

2-3  Disconnect the coupler under the oil filter.

2-4  Then add oil.

Change Oil.

- We recommend to change oil and oil filter every 6 to 12 months.

1-1  Open right side panel of AccuTOF.

1-2  Vent evacuation system (Please do vent process)

1-3  Remove Drain cap of RP drain tube.

(refer to right drawing)

1-4  Prepare a waste oil container.

Next page

1-5  Open drain valve and drain oil into the waste oil container.

1-6  After draining oil, refill pump using the “Add Oil” procedure above.

Change oil filter.

1  Disconnect Oil filter box.

Disconnect here

Disconnect Oil return tubing

(Rear side of the oil filter)

[Behind oil-filter Box]

- Disassemble the oil filter box

Unscrew bolts (each corner, 4peces)

(Using 4mm hexagonal screw driver)

- Open the oil-filter box and take out old filters

- Clean up inside of filter box.

- Replace oil filters (two pieces).

Filter 1

(EDWARDS part # = A223-04-077)

(5pc/set)

Filter2

(EDWARDS part # = A223-04-199)

- After replacing filter assemble the filter box by reversing the disassembly process.

# Text boxes

Unscrew

Unscrew